| Fridge Identification | Hospital Name | Month | Year |
| --- | --- | --- | --- |

| Daily maintenance | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Record Temperature Range 2-6°C internal thermometer |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Digital Display |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Check temperature recording chart. Document reasons for any spikes or deviations |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Initials of staff member performing check |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Check fridge for blood to be returned to pathology provider |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |

| Weekly maintenance | 1 | 2 | 3 | 4 | 5 |  |
| --- | --- | --- | --- | --- | --- | --- |
| Change temperature chart |  |  |  |  |  | Weekly: remove old temperature chart and annotate both old and new chart with initials, hospital name, date & time. Replace with new chart aligned to correct day and time. Account for any spikes by writing on old chart. |

| Monthly maintenance | Date | Pass (√ )Fail (X ) | Initials |
| --- | --- | --- | --- |
| Alarm back-up battery |  |  |  |
| *High Alarm activation (set at 5.5°C) |  |  |  |
| *Low alarm activation (set at 2.5°C) |  |  |  |
| Cleaning of internal surfaces and door seals |  |  |  |
| Test all alarms |  |  |  |
| Audible |  |  |  |
| Visible |  |  |  |
| Power Loss |  |  |  |
| Remote Alarm |  |  |  |

*AS3864.2-2012 recommends that this test is performed 6 monthly